## IS 13779_0_0_2020_0_TestReport
Annex-C
IS Number | Part | Section | Year | Number of Amendments | Title of the Standard | Type of Test | Qualitative or Quantitative | Clause/ Table/ Annex Reference | Test Parameter Name | To be tested and / reported by Lab (YES / NO) | Mandatory or Optional | Requirements (as per Indian Standard) - No specified values to be mentioned here | Test Result/ Observation | Test Request Requirements | Test Method Reference | Test Method Reference Standard | Sl No. under Clause/ Table/ Annex Reference | Sample Quantity | Equipment |  | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Clearance Distance for Indoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Clearance Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Creepage Distance for Indoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 100 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 150 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 300 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Creepage Distance for Outdoor Meter for Phase to Earth Voltage not exceeding 600 V | Specified limits for Active Power Consumption | Specified limits for Active Power Consumption | Specified limits for Apprent Power Consumption | Specified limits for Apprent Power Consumption | Specified limits for Apprent Power Consumption IN Current Circuit - Direct connection single phase and polyphase (per phase) for Class 1 | Specified limits for Apprent Power Consumption IN Current Circuit - Direct connection single phase and polyphase (per phase) for Class 1 | Specified limits for Apprent Power Consumption IN Current Circuit - Direct connection single phase and polyphase (per phase) for Class 2 | Specified limits for Apprent Power Consumption IN Current Circuit - Direct connection single phase and polyphase (per phase) for Class 2 | Specified limits for Apprent Power Consumption IN Current Circuit - Connection through current transformers Single phase and polyphase (per phase) for Class 1 | Specified limits for Apprent Power Consumption IN Current Circuit - Connection through current transformers Single phase and polyphase (per phase) for Class 1 | Specified limits for Apprent Power Consumption IN Current Circuit - Connection through current transformers Single phase and polyphase (per phase) for Class 2 | Specified limits for Apprent Power Consumption IN Current Circuit - Connection through current transformers Single phase and polyphase (per phase) for Class 2 | Specified limits for Variation in Percentage Error due to Short time over currents for Direct Connection - Class 1 | Specified limits for Variation in Percentage Error due to Short time over currents for Direct Connection - Class 1 | Specified limits for Variation in Percentage Error due to Short time over currents for Direct Connection - Class 2 | Specified limits for Variation in Percentage Error due to Short time over currents for Direct Connection - Class 2 | Specified limits for Variation in Percentage Error due to Short time over currents for Connection through Current transformer - Class 1 | Specified limits for Variation in Percentage Error due to Short time over currents for Connection through Current transformer - Class 1 | Specified limits for Variation in Percentage Error due to Short time over currents for Connection through Current transformer - Class 2 | Specified limits for Variation in Percentage Error due to Short time over currents for Connection through Current transformer - Class 2 | Specified limits for Variation in Percentage Error due to Self Heating - Class 1 | Specified limits for Variation in Percentage Error due to Self Heating Class 1 | Specified limits for Variation in Percentage Error due to Self Heating at Power factor 1- Class 2 | Specified limits for Variation in Percentage Error due to Self Heating at Power factor 1- Class 2 | Specified limits for temperature rise at any point of the external surface of the meter | Specified limits for temperature rise at any point of the external surface of the meter | Specified Limits for Insulation Resistance | Secified Limits for Insulation Resistance | Specified Limits for Variation in Percentage Error due to Earth/Phase Fault- or Class 1 | Specified Limits for Variation in Percentage Error due to Earth/Phase Fault- or Class 1 | Specified Limits for Variation in Percentage Error due to Earth/Phase Fault- or Class 2 | Specified Limits for Variation in Percentage Error due to Earth/Phase Fault- or Class 2 | Specified Limits difference between the percentage error when the meter is carrying a single-phase load at basic current and unity power factor and the percentage error when the meter is carrying balance polyphase load at basic current and unity power factor - Class 1 | Specified Limits difference between the percentage error when the meter is carrying a single-phase load at basic current and unity power factor and the percentage error when the meter is carrying balance polyphase load at basic current and unity power factor - Class 1 | Specified Limits difference between the percentage error when the meter is carrying a single-phase load at basic current and unity power factor and the percentage error when the meter is carrying balance polyphase load at basic current and unity power factor - Class 2 | Specified Limits difference between the percentage error when the meter is carrying a single-phase load at basic current and unity power factor and the percentage error when the meter is carrying balance polyphase load at basic current and unity power factor - Class 2 | Specified Limits for Percentage Error Limits for Meters of Class 1 | Specified Limits for Percentage Error Limits for Meters of Class 1 | Specified Limits for Percentage Error Limits for Meters of Class 2 | Specified Limits for Percentage Error Limits for Meters of Class 2 | Specified Limits for Mean Temperature Coefficient for Meter of Class 1 | Specified Limits for Mean Temperature Coefficient for Meter of Class 1 | Specified Limits for Mean Temperature Coefficient for Meter of Class 2 | Specified Limits for Mean Temperature Coefficient for Meter of Class 2 | Specified Limits for Starting Current for Direct connection- Class 1 meter | Specified Limits for Starting Current for Direct connection- Class 1 meter | Specified Limits for Starting Current for Direct connection- Class 2 meter | Specified Limits for Starting Current for Direct connection- Class 2 meter | Specified Limits for Starting Current for Connection through current transformer.- Class 1 meter | Specified Limits for Starting Current for Connection through current transformer.- Class 1 meter | Specified Limits for Starting Current for Connection through current transformer.- Class 2 meter | Specified Limits for Starting Current for Connection through current transformer.- Class 2 meter | Specified Limits for Repeatability of Error Test -Class 1 | Specified Limits for Repeatability of Error Test -Class 1 | Specified Limits for Repeatability of Error Test -Class 2 | Specified Limits for Repeatability of Error Test -Class 2 | Specified Limits for Variation in Percentage error at accuracy class index at Ib | Specified Limits for Variation in Percentage error at accuracy class index at Ib | Specified Limits for Variation in Percentage error at accuracy class index at 0.05 Ib | Specified Limits for Variation in Percentage error at accuracy class index at 0.05 Ib | Specified Limits for Variation in Percentage error at accuracy class index at Imax | Specified Limits for Variation in Percentage error at accuracy class index at Imax | Specified Limits for Variation in Percentage error at accuracy class index at Ib | Specified Limits for Variation in Percentage error at accuracy class index at Ib | Specified Limits for Variation in Percentage error at accuracy class index at 0.05 Ib | Specified Limits for Variation in Percentage error at accuracy class index at 0.05 Ib | Specified Limits for Variation in Percentage error at accuracy class index at Imax | Specified Limits for Variation in Percentage error at accuracy class index at Imax | Secified Limits for Insulation Resistance | Secified Limits for Insulation Resistance | Specified Limits for Insulation Resistance between circuits not intended to be connected together in service | If the test is quantitative, the Min and Max value to be written after applying the tolerance given in the standard. If only upper and lower limit is given, mention ONLY Min or ONLY Max balue and leave the other cell blank. If the test is Qualitative, both Only Min column to be filled by the qualitiative requirement, the Max column to be left blank
THIS ROW TO BE LEFT BLANK | [20 empty] | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min | Max | Min
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | General | Yes | Mandatory | Meters shall be designed and constructed in such a way as to avoid introducing any danger in normal use and under normal conditions, so as to ensure especially: a) Personal safety against electric shock; b) Personal safety against effects of excessive temperature; c) Safety against spread of fire; and d) Protection against solid objects, dust and water | Satisfactory/Unsatisfactory - Danger in normal use observed |  |  | IS 13779 | To be only filled if the Column I and Column J are having same information and Column M is having different information | NOT TO BE FILLED
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | General | Yes | Mandatory | All parts which are subject to corrosion under normal conditions shall be effectively protected | Satisfactory/Unsatisfactory-Parts subjected to corossion are not effectively protected/ Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | General | Yes | Mandatory | Any protective coating shall not be liable to damage by ordinary handling nor damage due to exposure to air, used under normal working conditions | Satisfactory/Unsatisfactory - Damage observed in Protective coating by ordinary handling, exposure to air, used under normal working conditions/ Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.2 | Meter Case | Yes | Mandatory | The meter shall have a reasonably dust/moisture-proof case, which shall be sealed by the manufacturer in such a way that the internal parts of the meter are accessible only after breaking such distinctive seal(s) | Satisfactory/Unsatisfactory - Internal parts of the meter are accessible without breaking distinctive seal(s) |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.2 | Meter Case | Yes | Mandatory | The cover shall not be removable without the use of a tool. | Satisfactory/Unsatisfactory - Cover is removable without use of tool |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.2 | Meter Case | Yes | Mandatory | The case shall be so constructed and arranged that any non-permanent deformation cannot prevent the satisfactory operation of the meter | Satisfactory/Unsatisfactory - Non Permanent deformation prevents the satisfactory operation of the meter |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.2 | Meter Case | Yes | Mandatory | Unless otherwise specified, the meters having a case wholly or partially made of metal, shall be provided with a protective earth terminal. | Satisfactory/Unsatisfactory - Protective eahing terminal not provided |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.3 | Window | Yes | Mandatory | If the cover is not transparent one or more windows shall be provided for reading the displays and observation of the operation indicator, if fitted. | Satisfactory/Unsatisfactory- Windows for reading the display not provided and cover is not transperent/ Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.3 | Window | Yes | Mandatory | These windows shall be covered by toughened transparent material which cannot be removed undamaged without breaking the seal(s). | Satisfactory/Unsatisfactory- Toughened transperent material coverng the window can be removed without breaking the seal/ Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | The material of which the terminal block is made shall be capable of passing the tests given in IS 13360 (Part 6/Sec 17) for a temperature of 135 °C and a pressure of 1.8 MPa (Method A) | Please refer Test report of IS 13360(Part 6/Sec 17) |  |  | IS 13360(Part 6/Sec 17) |  |  |  | Test report format to be prepared for IS 13360(Part 6/Sec 17)
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | The holes in the insulating material which form an extension of the terminal holes shall be of sufficient size to also accommodate the insulation of the conductors | Satisfactory/Unsatisfactory-Holes are not off sufficient size to accommodate insulation of conductors/ Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | The manner of fixing the conductors to the terminals shall ensure adequate and durable contact such that there is no risk of loosening or undue heating. | Satisfactory/Unsatisfactory - Adequate and Durable contact not ensured |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | Screw connections transmitting contact force and screw fixings which may be loosened and tightened several times during the life of the meter shall screw into a metal nut | Satisfactory/Unsatisfactory - Screw fixing did not screw into a metal nut |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | All parts of each terminal shall be such that the risk of corrosion resulting from contact with any other metal part is minimized. | Satisfactory/Unsatisfactory- Risk of corossion is observed |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | The current circuit connections can be terminated or guided through the measuring element of meter. The cable/conductor in such a design shall be held firmly in position by suitable guiding and fixing arrangement and that contact pressure is not transmitted through insulating material. For current circuits, the voltage is considered to be the same as for the related voltage circuit | Satisfactory/Unsatisfactory - Contact pressure transmission throgh insulating material observed |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | The voltage circuit connection can be directly terminated on a suitable terminal arrangement or tapped by suitable arrangement which shall not cause burnouts and corrosion due to outdoor usage | Satisfactory/Unsatisfactory-Voltage circuit connection is not directly terminated nor tapped by suitabel arrangement |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | Terminals with different potentials which are grouped close together shall be protected against accidental short-circuiting | Satisfactory/Unsatisfactory - Terminals with different potentials which are grouped close together are not protected against accidental short-circuiting |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | The terminals, the conductor fixing screws, or the external or internal conductors shall not be liable to come in to contact with terminal covers | Satisfactory/Unsatisfactory - The terminals, the conductor fixing screws, or the external or internal conductors came in to contact with terminal covers |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | The protective earth terminal, if any: a) Shall be electrically bonded to the accessible metal parts; b) Should, if possible, form part of meter base; c) Should preferably be located adjacent to its terminal block; d) Shall accommodate a conductor having a cross section at least equivalent to the main current conductors but with a lower limit of 6 mm2 and an upper limit of 16 mm2 (for copper conductors); and e) Shall be clearly identified by the earthing symbol | Satisfactory/Unsatisfactory - Did not accommodate a conductor having a cross section at least equivalent to the main current conductors (from 6mm2 to 16 mm2)Not electrically bonded to accessible metal parts, Earthing symbol not marked /Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.4 | Terminals, Terminals Block(s) and Protective Earth Terminal | Yes | Mandatory | After installation, it shall not be possible to loosen the protective earth terminal without the use of a tool. | Satisfactory/Unsatisfactory - possible to loosen the protective earth terminal without the use of a tool/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.5.1 | Short Terminal Cover | Yes | Mandatory | The terminals, their fixing screws and the insulated compartment housing them shall be enclosed by a cover with a provision for sealing | Satisfactory/Unsatisfactory- The terminals, their fixing screws and the insulated compartment housing them are not be enclosed by a cover with a provision for sealing/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.5.2 | Extended Terminal Cover | Yes | Mandatory | The terminals, their fixing screws, a suitable length of external insulated conductor and its insulation shall be enclosed by a cover with a provision for sealing | Satisfactory/Unsatisfactory - The terminals, their fixing screws, a suitable length of external insulated conductor and its insulation are not enclosed by a cover with a provision for sealing/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.5.2 | Extended Terminal Cover | Yes | Mandatory | The wiring with this type of cover shall be carried out from the rear of the meter board | Satisfactory/Unsatisfactory - The wiring is not carriedout from rear of the meter board/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.5.2 | Extended Terminal Cover | Yes | Mandatory | The fixing screws used on the terminal cover for fixing and sealing in 6.5.1 and 6.5.2 shall be held captive in the terminal cover | Satisfactory/Unsatisfactory - The fixing screws are not held in terminal cover/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.5.2 | Extended Terminal Cover | Yes | Mandatory | When the meter is mounted on the meter board, no access to the terminals shall be possible without breaking seal(s) of the terminal cover | Satisfactory/Unsatisfactory - Access to the terminals without breaking seal(s) of the terminal cover observed/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.6 | Clearance Distance | Yes | Mandatory | Clearance Distance. mm Min | Number of Decimal Places - Value upto one decimal places | Phase-to-Earth Derived from Rated System Voltages; Indoor/Outdoor meter |  | IS 13779 |  |  |  |  | 0.8 |  | 1.5 |  | 3 |  | 5.5 |  | 1.5 |  | 3 |  | 5.5 |  | 8
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.6 | Creepage Distance | Yes | Mandatory | Creepage Distance. mm Min | Number of Decimal Places - Value upto one decimal places | Phase-to-Earth Derived from Rated System Voltages; Indoor/Outdoor meter |  | IS 13779 | [20 empty] | 1.4 |  | 1.6 |  | 3.2 |  | 6.3 |  | 3.2 |  | 5 |  | 10 |  | 20
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.7 | Insulating Encased Meter | Yes | Mandatory | A meter having a durable and substantially continuous enclosure made wholly of insulating material, including the terminal cover, which envelops all metal parts with the exception of small parts, for example, name-plate screws, suspensions and rivets. If such small parts reaccessible by the standard test finger (see IS 1401) from outside the case, then they shall be additionally isolated from live parts by supplementary insulation against failure of basic insulation or loosening of live parts. | Satisfactory/Unsatisfactory - Small parts reaccessible by the standard test finger from outside the case, are not additionally isolated from live parts by supplementary insulation /Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.8 | Resistance to heat and fire | Yes | Mandatory | They should not be ignited by thermic overload of live parts in contact with them | Satisfactory/Unsatisfactory - Ignited by thermic overload of live parts |  | 12.4 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.9 | Protection against Penetration of Dust and Water | Yes | Mandatory | The meter shall conform to the degree of protection as given below: IP 51 for indoor meter and IP 54 for outdoor meter | Satisfactory/Unsatisfactory - Ingress of Dust in sufficient quantity to prevent operation of machine, Harmful effect of dripping water observed - Indoor meter; Unsatisfactory -Ingress of Dust in sufficient quantity to prevent operation of machine, Harmful effect by Water projected by a nozzle against the machine from any direction observed/Not Applicable | Indoor/Outdoor meter |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | Display of Measured Values | Yes | Mandatory | In case of an electronic display, the corresponding non-volatile memory shall have a minimum retention time of 5 years. Manufacturer shall declare the retention time. | To be checked on Delcaration basis | Manufacturer Declaration on retention time |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | Display of Measured Values | Yes | Mandatory | In the case of multiple values presented by a single display, it must be possible to display the contents of all relevant memories. | Satisfactory/Unsatisfactory-Contents of all relevant memories not displayed/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | Display of Measured Values | Yes | Mandatory | When displaying the memory, identification of each tariff/parameter shall be available | Satisfactory/Unsatisfactory - Identification of each tariff/parameter not available/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | Display of Measured Values | Yes | Mandatory | The current tariff shall be indicated. | Satisfactory/Unsatisfactory-Current tariff not indicated/Not Applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | Display of Measured Values | Yes | Mandatory | The principal unit for the measured values shall be kilo watthour (kWh) or the mega watthour (MWh). | kWh/MWh/Unsatisfactory-Principal unit for measured values is neither kilo watthour (kWh) nor the mega watthour (MWh). |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.1 | Display of Measured Values | Yes | Mandatory | The register shall be able to record and display starting from zero, for a minimum of 1 500 h, the energy corresponding to rated maximum current at reference voltage and unity power-factor. Register should not rollover in between this duration. | Satisfactory/Unsatisfactory-Rollover occurred |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.11 | Output Device | Yes | Mandatory | The meter shall have a test output accessible from the front and be capable of being monitored with suitable testing equipment | Satisfactory/Unsatisfactory-Test output not accessible from front and cant be monitored |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.11 | Output Device | Yes | Mandatory | Since the sequence of test output pulses may not be homogeneous, the manufacturer shall state the minimum number of pulse counts necessary to ensure that the error contribution due to such non-homogeneity does not exceed 1/10th of the specified error limits at various points and consistent with desired resolution | Satisfactory/Unsatisfactory-The minimum number of pulse counts necessary to ensure thatthe error contribution doesnor exceed 1/10th of the specifed error limits not stated |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.11 | Output Device | Yes | Mandatory | The resolution of the test output in the form of pulses or high resolution register, whether accessible on the meter through external display, shall be sufficient to conduct satisfactorily accuracy test at the lowest load in less than 5 min and starting current test in less than 10 min. | Satisfactory/Unsatisfactory - The resolution of the test output in the form of pulses or high resolution register, is not sufficient to conduct satisfactorily accuracy test at the lowest load in less than 5 min and starting current test in less than 10 min. |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 6.11 | Output Device | Yes | Mandatory | Additional high resolution register with minimum decimal of two digits shall also be available in the meter for testing purpose as mentioned in this standard | Satisfactory/Unsatisfactory - Additional high resolution register with minimum decimal of two digits not provided |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | Every meter shall bear the following information - Manufacturer’s name and/or trade mark and the place of manufacture | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | Designation of type (see 3.1.4) and, if required, space for approval mark | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The number of phases and the number of wires for which the meter is suitable (for example, single phase 2 wire, three phase 3 wire, three phase 4 wire). These markings may be replaced by the graphical symbols (as per IS 12032 series); | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The serial number and year of manufacture. If the serial number is marked on a plate fixed to the cover, the number is to be marked also on the meter base or displayed | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The reference voltage in one of the following forms: The number of elements if more than one, and the voltage at the meter terminals of the voltage circuit(s). The nominal voltage of the system or the secondary voltage of the instrument transformer to which the meter is intended to be connected. | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | Principal unit in which the meter reads, for example, kWh | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The rated secondary current of the transformer(s) to which the meter should be connected, for example, –/5 A: the basic current and the rated maximum current of the meter may be included in the type designation | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The reference frequency in Hz | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The meter constant, for example in the form; X Wh/imp or X imp/kWh; | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The class index of the meter | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The reference temperature if different from 27 °C; | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The sign of double square for insulating encased meters; and | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | Country of manufacture | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The information (d) to (k) shall be marked on a name plate preferably placed within the meter. | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | The marking shall be indelible, distinct and readable from outside the meter | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | If the meter registers energy through instrument transformer(s), which are accounted in the meter constant, the transformation ratio(s) shall be marked | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Mandatory | Standard symbols shall be used | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.1 | Marking | Yes | Optional | The product may be marked with the Standard Mark | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.2 | Connections, Diagrams and Terminal Marking | Yes | Mandatory | Every meter shall be indelibly marked with a diagram of connections | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.2 | Connections, Diagrams and Terminal Marking | Yes | Mandatory | For poly-phase meters, this diagram shall also show the phase sequence for which the meter is intended. | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.2 | Connections, Diagrams and Terminal Marking | Yes | Mandatory | If the meter terminals are marked, this marking shall appear on the diagram. | Satisfactory/Unsatisfactory - Not marked |  |  | IS 13779 |  |  |  | Not to be masked by sample cell
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.2 | Connections, Diagrams and Terminal Marking | Yes | Mandatory | When a number of meters are connected to single distributing mains for registering electricity supplied to different consumer loads, separate service lines-phase(s) and neutral, shall be used for each meter | Satisfactory/Unsatisfactory-Separate service lines-phase(s) and neutral not used/Not applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 7.2 | Connections, Diagrams and Terminal Marking | Yes | Mandatory | Each independently metered consumer load must be directly connected to distributing mains through its meter connected in specified phase sequence so as to meet accuracy requirements of this standard. | Satisfactory/Unatisfactory/Not applicable |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 8.1 | Temperature Range | Yes | Mandatory |  |  |  |  | IS 13779 |  |  |  | Only requirement is specified. No test method is specified in IS. Comment sent to ETD
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 8.2 | Relative Humidity | Yes | Mandatory |  |  |  |  | IS 13779 |  |  |  | Only requirement is specified. No test method is specified in IS. Comment sent to ETD
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.1.1, Table 9 | Power Consumption in Voltage Circuit | Yes | Mandatory | Active Power Consumption.W. Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [37 empty] | 1.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.1.1, Table 9 | Power Consumption in Voltage Circuit | Yes | Mandatory | Apparent Power Consumption. VA. Max | To be reported as Whole number |  |  | IS 13779 | [39 empty] | 10
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.1.2, Table 10 | Power Consumption in Current Circuit | Yes | Mandatory | Apparent Power Consumption in current circuit. VA. Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [41 empty] | 4 |  | 2.5 |  | 4 |  | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 9.2.2 | Voltage Dips and Interruptions | Yes | Mandatory | Voltage dips and interruptions shall not produce a change in the register of more than 0.01 kWh and the test output shall not produce a signal equivalent to more than 0.01 kWh. | Satisfactory/Unsatisfactory - Change in the register of more than 0.01 kWh, the test output produced a signal equivalent to more than 0.01 kWh. |  | 12.7.2 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 9.2.2 | Voltage Dips and Interruptions | Yes | Mandatory | When the voltage is restored, the meter shall not have suffered degradation of its metrological characteristics. | Satisfactory/Unsatisfactory - The meter suffered degradation of its metrological characteristics |  | 12.7.2 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.2.3, Table 12 | Variation due to Short time Over currents | Yes | Mandatory | Limits of Variation in Percentage Error for Variation due to Short time over currents. Percentage. Max | Number of Decimal Places - Value upto one decimal places |  | 12.7.3 | IS 13779 | [49 empty] | 1.5 |  | 1.5 |  | 0.5 |  | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.3 | Influence of Self Heating | Yes | Mandatory | Limits of Variation in Percentage Error for Influence of Self heating at I max, Unity Powerfactor. Percentage. Max | Number of Decimal Places - Value upto one decimal places |  | 12.7.4 | IS 13779 | [57 empty] | 0.7 |  | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.3 | Influence of Self Heating | Yes | Mandatory | Limits of Variation in Percentage Error for Influence of Self heating at Imax, 0.5 lagging powerfactor. Percentage. Max | Number of Decimal Places - Value upto one decimal places |  | 12.7.4 | IS 13779 | [57 empty] | 1 |  | 1.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.4 | Influence of Heating | Yes | Mandatory | The temperature rise at any point of the external surface of the meter. °C. Max | To be reported as Whole number |  | 12.7.5 | IS 13779 | [61 empty] | 20
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 9.4 | Influence of Heating | Yes | Mandatory | Visual inspection for deformation, if any of the meter including terminal blocks is to be carried out. There shall be no deformation | Satisfactory/Unsatisfactory-Deformation observed |  | 12.7.5 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.4 | Influence of Heating | Yes | Mandatory | Insulation Resistance after Influence of heating. Between, on the one hand,all the current and voltage circuits as well as the auxiliary circuits whose reference voltage is over 40V, connected together, and,on the other hand, earth. Mega Ohm. Min | To be reported as Whole number |  | 12.7.5 | IS 13779 | [62 empty] | 5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.4 | Influence of Heating | Yes | Mandatory | Insulation Resistance after Influence of heating. Between, Between circuits not intended to be connected together in service. Mega Ohm. Min | To be reported as Whole number |  | 12.7.5 | IS 13779 | [62 empty] | 50
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 9.5 | Insulation | Yes | Mandatory | Impulse Voltage Test | Satisfactory/Unsatisfactory- Did not withstood the test/Not Applicable |  | 12.7.6 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 9.5 | Insulation | Yes | Mandatory | AC Voltage Test - Type/Acceptance test | Satisfactory/Unsatisfactory- Flashover or Disruptive Dischagre or Puncture has occured |  | 12.7.6 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.5 | Insulation | Yes | Mandatory | Insulation Resistance. Between, on the one hand,all the current and voltage circuits as well as the auxiliary circuits whose reference voltage is over 40V, connected together, and,on the other hand, earth. Mega Ohm. Min | Mega Ohm. To be reported as Integer |  | 12.7.6 | IS 13779 | [62 empty] | 5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.5 | Insulation | Yes | Mandatory | Insulation Resistance. Between, Between circuits not intended to be connected together in service. Mega Ohm. Min | Mega Ohm. To be reported as Integer |  | 12.7.6 | IS 13779 | [62 empty] | 50
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 9.6 | Immunity to Earth/Phase Fault | Yes | Mandatory | After the test, the meter shall show no damage and shall operate correctly | Satisfactory/Unsatisfactory - Damage occurred, The meter did not operate properly/Not applicable |  | 12.8 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 9.6 | Immunity to Earth/Phase Fault | Yes | Mandatory | Change of Error due to Earth/Phase Fault- Limits of Variation in Percentage Error. Percentage. Max | Number of Decimal Places - Value upto one decimal places |  | 12.8 | IS 13779 | [65 empty] | 0.7 |  | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 10.1 | Test for Immunity to Electrostatic Discharge | Yes | Mandatory | The application of the electrostatic discharge shall not produce a change in the cumulative register of more than 0.01 kWh and the test output shall not produce a signal equivalent to more than 0.01 kWh. | Satisfactory/Unsatisfactory - change in the cumulative register of more than 0.01 kWh observed, The test output produced a signal equivalent to more than 0.01 kWh |  | 12.9.2 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 10.1 | Test for Immunity to Electromagnetic HF Fields | Yes | Mandatory | The application of the electrostatic discharge shall not produce a change in the cumulative register of more than 0.01 kWh and the test output shall not produce a signal equivalent to more than 0.01 kWh. | Satisfactory/Unsatisfactory - change in the cumulative register of more than 0.01 kWh observed, The test output produced a signal equivalent to more than 0.01 kWh |  | 12.9.3 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 10.1 | Fast Transient Burst Test | Yes | Mandatory | Refer requirements for Row 114 to 128 |  |  | 12.9.4 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 10.1 | Surge Immunity test | Yes | Mandatory | The application of test shall not produce change in register of more than 0.01 kWh and the test output shall not produce a signal equivalent to more than 0.01 kWh. | Satisfactory/Unsatisfactory - change in the cumulative register of more than 0.01 kWh observed, The test output produced a signal equivalent to more than 0.01 kWh |  | 12.9.5 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 10.2 | Radio Interference Suppression | Yes | Mandatory | The meter shall not generate conducted or radiated noise which could interfere with other equipment. | Satisfactory/Unsatisfactory-Generate conducted or radiated noise which interfered with other equipment. |  | 12.9.6 | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | The difference between the percentage error when the meter is carrying a single-phase load at basic current and unity power factor and the percentage error when the meter is carrying balance polyphase load at basic current and unity power factor. Percentage. Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [69 empty] | 1.5 |  | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.05 Ib ≤ I < 0.1 Ib at Unity Power factor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1.5 | 1.5 | 2.5 | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.1 Ib ≤ I ≤ Imax at Unity Power factor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1 | 1 | 2 | 2
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.1 Ib ≤ I < 0.2 Ib at 0.5 lagging Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1.5 | 1.5 | 2.5 | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.1 Ib ≤ I < 0.2 Ib at 0.8 leading Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1.5 | 1.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.2 Ib ≤ I ≤ Imax at 0.5 lagging Power factor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1 | 1 | 2 | 2
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.2 Ib ≤ I ≤ Imax at 0.8 leading Power factor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1 | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Optional | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.2 Ib to Ib at 0.25 lagging Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 3.5 | 3.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Optional | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Direct Connected Meters 0.2 Ib to Ib at 0.5 leading Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 2.5 | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.02 Ib ≤ I < 0.05 Ib at unity Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1.5 | 1.5 | 2.5 | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.05 Ib ≤ I ≤ Imax at unity Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1 | 1 | 2 | 2
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.05 Ib ≤ I < 0.1 Ib at 0.5 lagging Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1.5 | 1.5 | 2.5 | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.05 Ib ≤ I < 0.1 Ib at 0.8 leading Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1.5 | 1.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.1 Ib ≤ I ≤ Imax at 0.5 lagging Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1 | 1 | 2 | 2
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.1 Ib ≤ I ≤ Imax at 0.8 leading Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 1 | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Optional | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.1 Ib to Ib at 0.25 lagging Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 3.5 | 3.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Optional | Percentage Error Limits for Single-Phase Meters and Polyphase Meters with Balanced Loads for Trasformer Operated Meters 0.1 Ib to Ib at 0.5 leading Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 2.5 | 2.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Polyphase Meters Carrying a Single Phase Load, But with Balanced Polyphase Voltages Applied to Voltage Circuits for Direct Connected Meters 0.1 Ib ≤ I ≤ Imax at unity Power factor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 2 | 2 | 3 | 3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Polyphase Meters Carrying a Single Phase Load, But with Balanced Polyphase Voltages Applied to Voltage Circuits for Direct Connected Meters 0.2 Ib ≤ I ≤ Imax at 0.5 lagging Power factor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 2 | 2 | 3 | 3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Polyphase Meters Carrying a Single Phase Load, But with Balanced Polyphase Voltages Applied to Voltage Circuits for Trasformer Operated Meters 0.05 Ib ≤ I ≤ Imax at unity Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 2 | 2 | 3 | 3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.1 | Limits of Error Due to Variation of the Current | Yes | Mandatory | Percentage Error Limits for Polyphase Meters Carrying a Single Phase Load, But with Balanced Polyphase Voltages Applied to Voltage Circuits for Trasformer Operated Meters 0.1 Ib ≤ I ≤ Imax at 0.5 lagging Powerfactor. Percentage. Min and Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [72 empty] | 2 | 2 | 3 | 3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Voltage variation ±10 percent, at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 0.7 |  | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Voltage variation ±10 percent, at Ib Current, at 0.5 lagging Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 1 |  | 1.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Frequency variation ± 5 percent, at Ib Current, at unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 0.8 |  | 1.3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Frequency variation ± 5 percent, at Ib Current, at 0.5 lagging Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 1 |  | 1.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Wave form: 10 percent of third harmonic in the current, t Ib Current, at unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 0.6 |  | 0.8
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Reversed phase sequence, at 0.1 Ib, at unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 1.5 |  | 1.5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Voltage Unbalance, at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 2 |  | 4
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | d.c. and even harmonics in AC current circuit, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 3 |  | 6
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Continuous magnetic induction of external origin, at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 2 |  | 3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Magnetic induction of external origin 0.5 mT, at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 2 |  | 3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Electromagnetic HF fields, at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 2 |  | 3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Operation of accessories, at 0.05 Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 0.5 |  | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Continuous ‘abnormal’ magnetic induction of external origin at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 4 |  | 4
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | ‘Abnormal’ ac magnetic induction of external origin 10 mT at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 4 |  | 4
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.2 | Influence Quantities | Yes | Mandatory | Electrical fast transient burst at Ib Current, at Unity Powerfactor. Percentage. Max. | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [73 empty] | 4 |  | 6
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.3 | Limits of Error Due to Ambient Temperature Variation | Yes | Mandatory | Temperature Coefficient - From 0.1 Ib to Imax at Unity Powerfactor. Percentage. Max | Number of Decimal Places - Value upto Two decimal places |  |  | IS 13779 | [77 empty] | 0.05 |  | 0.1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.3 | Limits of Error Due to Ambient Temperature Variation | Yes | Mandatory | Temperature Coefficient - From 0.2 Ib to Imax at 0.5 lagging Powerfactor. Percentage. Max | Number of Decimal Places - Value upto Two decimal places |  |  | IS 13779 | [77 empty] | 0.07 |  | 0.15
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 11.4.1 | Initial Start-up of the Meter | Yes | Mandatory | The meter shall be functional within five seconds after the rated voltage is applied to meter terminals | Satisfactory/Unsatisfactory - The meter did not function within five seconds after the rated voltage is applied to meter terminals |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 11.4.2 | Running with No-Load | Yes | Mandatory | When the voltage is applied with no current flowing in the current circuit, the test output of the meter shall not produce more than one output pulse count | Satisfactory/Unsatisfactory - When the voltage is applied with no current flowing in the current circuit, the test output of the meter produced more than one output pulse count |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.5 | Starting Current | Yes | Mandatory | Starting Current. Percentage.Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [81 empty] | 0.4 |  | 0.5 |  | 0.2 |  | 0.3
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 11.6 | Meter Constant | Yes | Mandatory | Relation between the test output and the indication in the display shall comply with the marking on the name plate | Satisfactory/Unsatisfactory - Relation between the test output and the indication in the display didnot comply with the marking on the name plate |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 11.7 | Repeatability of Error Test | Yes | Mandatory | Repeatability of Error. Percentage. Max | Number of Decimal Places - Value upto one decimal places |  |  | IS 13779 | [89 empty] | 0.5 |  | 1
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 12.3 | Shock Test | Yes | Mandatory | After the test, the meter shall show no damage or change in cumulative kWh register. | Satisfactory/Unsatisfactory - Damage observed, Change in cumulative kWh register observed |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 12.3.1 | Shock Test | Yes | Mandatory | Variation in Percentage Error | Percentage. To be reported as Whole Number |  |  | IS 13779 | [93 empty] | 50 |  | 50 |  | 50
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 12.3.2 | Vibration Test | Yes | Mandatory | After the test, the meter shall show no damage or change in cumulative kWh register. | Satisfactory/Unsatisfactory - Damage observed, Change in cumulative kWh register observed |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 12.3.2 | Vibration Test | Yes | Mandatory | Variation in Percentage Error. Percentage. Max | To be reported as Whole Number |  |  | IS 13779 | [99 empty] | 50 |  | 50 |  | 50
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 12.3.3 | Spring Hammer Test | Yes | Mandatory | Terminal cover shall not sustain damage which could affect the function of the meter and it is not possible to touch live parts | Satisfactory/Unsatisfactory - Damage which affects the function of the meter observed, Live parts have become accessible |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 12.6.1 | Dry Heat Test | Yes | Mandatory | The meter shall show no damage or change in cumulative kWh register. These tests should not affect the functioning of the meters | Satisfactory/Unsatisfactory - Damage observed, Change in cumulative kWh register observed |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 12.6.1 | Cold Test | Yes | Mandatory | The meter shall show no damage or change in cumulative kWh register. These tests should not affect the functioning of the meters | Satisfactory/Unsatisfactory - Damage observed, Change in cumulative kWh register observed |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 12.6.1 | Damp Heat Cycle Test | Yes | Mandatory | The meter shall show no damage or change in cumulative kWh register. These tests should not affect the functioning of the meters | Satisfactory/Unsatisfactory - Damage observed, Change in cumulative kWh register observed |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Qualitative | 12.6.1 | Damp Heat Cycle Test | Yes | Mandatory | No trace of corrosion likely to affect the functional properties of the meter shall be apparent. | Satisfactory/Unsatisfactory -Corrosion likely to affect the functional properties of the meter obswerved |  |  | IS 13779
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 12.6.1 | Damp Heat Cycle Test | Yes | Mandatory | Insulation Resistance - between, on the one hand, all the current and voltage circuits as well as the auxiliary circuits whose reference voltage is over 40 V. Mega Ohm. Min | To be reported as Integer |  |  | IS 13779 | [105 empty] | 5
13779 |  |  | 2020 |  | A.C. STATIC WATTHOUR METERS, CLASS 1 AND 2 | Electrical | Quantitative | 12.6.1 | Damp Heat Cycle Test | Yes | Mandatory | Insulation Resistance - Insulation Resistance between circuits not intended to be connected together in service. Mega Ohm Min | To be reported as Integer |  |  | IS 13779 | [105 empty] | 50

## IS 13779_0_0_2020_0_TestRequest
Class of Meter | Phase-to-Earth Derived from Rated System Voltages | Retention Time | Connection
Class 1 | As declared | Manufacturer Declaration/Certificate | Direct Connection
Class 2 |  |  | Connection through Current Transformer